## Field data sheet
STATIONARY SOURCE MONITORING
LOG SHEET FOR ISOKINETIC FLOW MEASUREMENT (S.P.M.)
DATE : | dd/mm/yyyy |  |  | BAROMETRIC PRESSURE (P) = |  |  | 758 | mm Hg | PITOT TUBE COEFFICIENT (Cp) = |  |  | 0.967
TIME : | hrs/min |  |  | STATIC PRESSURE (ps) (+/-) = |  | observed on inclined manometer by removal of one end to atmosphere /13.6 (density of Hg) | 15/13.6= (-) 1.10 | mm Hg | AREA OF NOZZLE (B) IN M 2 = |  | 1/8 inch dia | 1.5835e-05
STACK ATTACHED TO : |  |  |  | ABSOLUTE STACK PRESSURE(Ps) =(P- ps)= |  |  | 759.1 | mm Hg | COEFFICIENT 'K' = (33.5 * Cp) / (SQRT(Ps * 30)) = |  |  | 0.21466
TRAVERSE | MANOMETER | STACK TEMP. | VELOCITY | FLOW RATE | DRY GAS METER | AMBIENT | TEMPERATURE | PRESSURE | SAMPLING | SAMPLING | FLOW | FLOW
POINT NO. | READING | Ts | =K*(SQRT(a * b)) | =A * B * 60000 | VACUUM | TEMPERATURE | CORRECTION | CORRECTION | FLOW RATE | TIME
Thimble No. | DELTA P mmwc | DEG 0 K | m / sec | LPM | mm Hg | DEG 0 K | ( Tm / Ts ) | Ps / (Ps-Pm) | LPM | MIN | LITS | M3
| (a) | (b) | ( A ) | ( C ) | ( Pm ) | ( Tm ) | ( D ) | ( E ) | C * D * E
9 | 90 | 505 | 45.7644 | 43.4798 | 3 | 274.7 | 0.54396 | 1.004 | 23.7451 | 15 | 356.1771 | 0.35618
13 | 35 | 516 | 28.8483 | 27.4081 | 3 | 274.5 | 0.53198 | 1.004 | 14.6383 | 15 | 219.5752 | 0.21958
14 | 47 | 515 | 33.3975 | 31.7302 | 3 | 274.1 | 0.53223 | 1.004 | 16.9549 | 15 | 254.3233 | 0.25432
15 | 36 | 516 | 29.2575 | 27.7969 | 3 | 275.2 | 0.53333 | 1.004 | 14.8839 | 15 | 223.2578 | 0.22326
18 | 32 | 540 | 28.2184 | 26.8097 | 3 | 275 | 0.50926 | 1.004 | 13.7073 | 15 | 205.6092 | 0.20561
19 | 27 | 542 | 25.9682 | 24.6719 | 3 | 274.2 | 0.5059 | 1.004 | 12.5311 | 15 | 187.9669 | 0.18797
20 | 56 | 433 | 33.4271 | 31.7584 | 3 | 274.3 | 0.63349 | 1.004 | 20.1984 | 15 | 302.9756 | 0.30298
23 | 31 | 497 | 26.6453 | 25.3151 | 3 | 274 | 0.55131 | 1.004 | 14.0118 | 15 | 210.1769 | 0.21018
| Average | 522.3333 | 31.9091 | 30.3161 | 3 | 274.6167 |  |  | Total Vspm | 30 | 1446.9095 | 1.4469